2

 
[image: image1.jpg]



THE EURASIAN ECONOMIC COMMISSION
THE BOARD

DECISION
	May 23, 2023
	No. 65
	Moscow


On amending the Rules for Classification of Medical Devices Depending on the Potential Risk of Use and on canceling certain decisions of the Eurasian Economic Commission 
In accordance with paragraph 4, Article 4 of the Agreement on Unified Principles and Rules for the Circulation of Medical Devices (Medical Accessories and Medical Equipment) within the Eurasian Economic Union dated December 23, 2014 and paragraph 23 of Annex No. 2 to the Rules of Procedure of the Eurasian Economic Commission approved by Decision No. 98 of the Supreme Eurasian Economic Council dated December 23, 2014, the Eurasian Economic Commission's Board decided:
1. To make amendments to the Rules for Classification of Medical Devices Depending on the Potential Risk of Use approved by Decision No. 173 of the Eurasian Economic Commission's Board dated December 22, 2015 according to the Annex.
2. To state that paragraph 1 of this Decision shall not cover legal relations associated with the medical device authorization, when the application for the procedures for authorization and expert assessment procedures, amending marketing authorization application (including post factum notification) or for expert opinion review for an authorized medical device is filed by the applicant on or before the effective date of this Decision.
3. To cancel:
a) Decision No. 116 of the Eurasian Economic Commission's Board dated July 24, 2018 On the Criteria for Differentiation of Medical Device Components for the Purposes of the Medical Device Marketing Authorization;
b) Decision No. 123 of the Eurasian Economic Commission's Board dated July 24, 2018 On Criteria for Inclusion into the Same Marketing Authorization Certificate of Several Modifications of a Medical Device related to the Same Medical Device Type According to the Medical Device Classification Applied in the Eurasian Economic Union;
4. This Decision shall come into effect after 180 calendar days have elapsed from the date of its official publication.
	Chairman of the Board
of the Eurasian Economic Commission
	M. Myasnikovich


ANNEX
to Decision of the
of the Eurasian Economic Commission
No. 65 dated May 23, 2023
AMENDMENTS
to be made to the Rules for Classification of Medical Devices Depending on Potential Risk of Use
The Rules for Classification of Medical Devices Depending on Potential Risk of Use shall be revised as follows:
"APPROVED BY
Decision of the Eurasian Economic Commission's Board

No. 173 dated December 22, 2015
(as amended by Decision of the Eurasian Economic Commission's Board

No. 65 dated May 23, 2023)
RULES
for Classification of Medical Devices
Depending on Potential Risk of Use
I. General provisions
1. These Rules are developed in order to implement the Agreement on Unified Principles and Rules for the Circulation of Medical Devices (Medical Accessories and Medical Equipment) within the Eurasian Economic Union dated December 23, 2014 and establish the procedure for classifying medical devices depending on the potential risk of harm such as injury or damage to the health of patients due to the use of medical devices, personnel operating medical devices, other persons, public health or the environment. 
2. These Rules shall apply to medical devices marketed within the Eurasian Economic Union.
3. For the purposes of these Rules, the concepts are used having the following meaning:
"active diagnostic medical devices" mean active medical devices designed to provide information for the purpose of diagnosis, monitoring of treatment or physiological state changes, disease state or congenital defects;
"active medical devices" mean medical devices used alone or in combination with other medical devices, the operation of which requires use of energy other than that produced by human, or the gravity. Medical devices intended for transfer of energy or substances from an active medical device to a patient without significant change are not active medical devices;
"active therapeutic medical devices" mean active medical devices intended to preserve, modify, replace or restore biological functions or structures associated with the treatment, alleviation of illness, injury or disability;
“analyte” means the sample component having measurable property;
"apheresis" means method of obtaining individual blood components, it is divided into plasmapheresis and cytopheresis;
"implantable medical devices" mean invasive medical devices, including partially or completely absorbable in the body, completely introduced into the human body or replacing the epithelial surface or the surface of the eye through surgical intervention and remaining at the introduction site after surgical procedure, as well as medical devices, partially introduced into the human body through surgical intervention and remaining at the introduction site after the surgical procedure for more than 30 days;
"invasive medical devices" mean the medical devices intended for full or partial introduction into the human body through its surface or through an aperture in the body; 
"artificial intelligence" mean a set of technological solutions allowing to simulate human cognitive functions (including self-learning and finding solutions without a predetermined algorithm), which includes information and communication infrastructure, software (including using machine learning methods), processes and services for data processing and search for solutions;
"medical devices for temporary use" mean medical devices intended for continuous use for a certain period (from 60 minutes to 30 days) in accordance with the instructions for use or operating instructions;
"medical devices for long-term use" mean medical devices intended for continuous use for more than 30 days in accordance with the instructions for use or operating instructions;
"medical devices for short-term use" mean medical devices intended for continuous use for no more than 60 minutes in accordance with the instructions for use or operating instructions; 
"medical device purpose" - mean the manufacturer's documented decision regarding the intended use of a medical device, based on its properties, reflected in the technical specifications, instructions for use or operating instructions;
"nanomaterial" means the material containing particles in unbound state, or particles in the form of aggregates or agglomerates, at least 50 percent of the particles of which have sizes within the range of 1 nm to 100 nm. At that, aggregates are understood to be particles consisting of fused or strongly bonded particles, 
and agglomerates are associations of weakly bonded particles. 
Nanomaterials also include graphene particles or carbon nanotubes with one or more outer dimensions less than 1 nm; 
"non-invasive medical devices" mean medical devices that are not intended for full or partial introduction into the human body through its surface or through body apertures; 
"non-professional user" means an individual who does not have special education in the relevant field. For an IVD medical device intended for use outside the laboratory environment, the user of such medical device shall be considered a non-professional user;
"body aperture" means any natural opening of the human body, as well as the eyeball outer surface or any permanent artificial opening (cavity);
"artificial intelligence technologies" mean technologies based on the artificial intelligence, including computer vision, natural language processing, speech recognition and synthesis, intelligent decision support and other promising areas;
"surgical invasive medical devices" mean invasive medical devices wholly or partially introduced into the human body through its surface or through body aperture by means of or in connection with surgical intervention. 
Other concepts used in these Rules shall be used in the meanings determined by the acts of the Eurasian Economic Union's governing bodies in the field of medical device marketing.
II. Classification of medical devices,
except for the IVD medical devices and software being the medical device
1. Classes of medical devices depending on potential risk of use
4. The medical devices depending on potential risk of use divide into 4 classes: 1, 2a, 2b and 3. 
Each medical device may be assigned to one class only.
The medical devices shall be classified on the basis of the following:
class 1 includes medical devices with low potential risk of use;
class 2a includes medical devices with medium potential risk of use;
class 2b includes medical devices with increased potential risk of use;
class 3 includes medical devices with high potential risk of use;
5. When classifying a medical device, the medical device purpose and conditions of its use, as well as the following criteria shall be considered:
a) duration of use of the medical device;
b) invasiveness of the medical device;
c) presence of contact of the medical device with the human body or the relation with it;
d) method of introducing the medical device into the human body (through a body aperture or surgically);
e) use of the medical device for vital organs and systems (heart, central circulation, central nervous system); 
f) use of energy sources.
2. Classification of non-invasive medical devices
6. Non-invasive medical devices belong to class 1, if paragraphs 7 through 9 of these Rules are not applied to them, with the exception of subparagraph "a" in paragraph 9 of these Rules. 
7. Non-invasive medical devices intended for storing organs, parts of organs or for storing or administrating blood, other liquids, gases, vapors or tissues into the human body belong to class 2a, including when they are used in combination with class 2a or higher active medical device.
8. Non-invasive medical devices designed to change the biological or physico-chemical composition and properties of blood, tissues, cells, other physiological fluids or liquids that shall enter the human body belong to class 2b. If their function is only particle filtration, centrifugation, gas or heat exchange, these medical devices are classified in class 2a.
9. Non-invasive medical devices that come into contact with damaged skin belong to: 
a) class 1, if they are used as mechanical barriers, for compression or exudate absorption;
b) class 2b, if they are used for wounds that can be cured only through secondary adhesion;
c) class 2a in other cases (including when medical devices are intended primarily to affect the wound microenvironment).
3. Classification of invasive medical devices
10. Invasive medical devices (except for surgical invasive ones), the use of which is associated with body apertures and which are not intended to be attached to an active medical device (except for class 1 active medical device), belong to:
class 1, if they are short-term application products; 
class 2a, if they are temporary medical devices. If these medical devices are temporarily used in the oral cavity up to the pharynx, in the ear canal up to the eardrum or in the nasal cavity, they belong to class 1; 
class 2b, if they are medical devices for long-term use. If these medical devices are used for long time in the oral cavity up to the pharynx, in the ear canal up to the eardrum or in the nasal cavity and cannot be absorbed by the mucous membrane, they belong to class 2a.
The invasive medical devices (except for the surgical invasive ones), use of which is associated with body apertures and which are intended to be attached to an active medical devices of class 2a or higher, belong to class 2a.
11. Surgical invasive medical devices for short-term use belong to class 2a, except for the following cases: 
a) if the medical devices are intended for diagnosis, observation, monitoring or correction of pathologies of the heart, central circulation or central nervous system and are in direct contact with organs or parts of these systems, they belong to class 3;
b) if the medical devices are reusable surgical instruments, they belong to class 1;
c) if the medical devices are designed to transfer energy in the form of ionizing radiation, they belong to class 2b;
d) if the medical devices are intended to cause biological effect or to dissolve completely or to significant extent, they belong to class 2b;
e) if the medical devices are intended for the administration of medicinal products by non-professional users, they belong to class 2b.
12. Surgical invasive medical devices for temporary use belong to class 2a, except for the following cases:
a) if the medical devices are intended for diagnosis, observation, monitoring or correction of pathologies of the heart or central circulation and are in direct contact with organs or parts of these systems, they belong to class 3;
b) if the medical devices are in direct contact with the central nervous system, they belong to class 3;
c) if the medical devices are designed to transfer energy in the form of ionizing radiation, they belong to class 2b;
d) if the medical devices are intended to cause biological effect or to dissolve completely or to significant extent, they belong to class 3;
e) if the medical devices undergo chemical changes in the human body, they belong to class 2b (except for medical devices intended for implantation in teeth or for administration of medicinal products).
13. Implantable medical devices, as well as surgical invasive medical devices for long-term use, belong to class 2b, except for the following cases:
a) if the medical devices are intended for implantation in teeth, they belong to class 2a;
b) if the medical devices are in direct contact with heart, central circulation or central nervous system, they belong to class 3;
c) if the medical devices are intended to cause biological effect or to dissolve completely or to significant extent, they belong to class 3;
d) if the medical devices undergo chemical changes in the human body, they belong to class 3 (except for the medical devices intended for implantation in teeth or for administration of medicinal products);

e) if the medical devices are active implantable medical devices, including implantable accessories for active implantable medical devices, they belong to class 3;
f) if the medical devices are breast implants, they belong to class 3;
g) if the medical devices are total or partial hip, knee or shoulder joint prostheses, they belong to class 3;
h) if the medical devices are intervertebral disc prostheses or implantable medical devices that come into contact with the spine, they belong to class 3.
4. Particulars of classification of active medical devices

14. Active medical devices are classified considering the following particulars:
a) active therapeutic medical devices intended for energy transfer or energy exchange belong to class 2a. If energy transfer to the human body or energy exchange with it poses potential hazard due to the medical device features, taking into account the nature, density and location of the energy impact on body parts, these medical devices belong to class 2b (including active medical devices designed to create ionizing radiation, radiation therapy);
b) active medical devices designed to monitor or control the class 2b active therapeutic medical devices belong to class 2b. If active medical devices are intended to monitor or control active implantable medical devices, the medical devices are classified in class 3.
15. Active diagnostic medical devices belong to class 2a if they are intended for:
a) transfer of energy absorbed by a human. If the medical device function is to illuminate the patient's body in the visible spectrum, such a medical device belongs to class 1;
b) concept of distribution of the radiopharmaceutical medicinal products introduced into the human body;
c) providing direct diagnostics or monitoring of vital body functions. If the medical devices are intended to monitor vital physiological parameters, changes in which could lead to immediate danger to a patient (e.g., change in function of the heart, respiration or central nervous system activity), they belong to class 2b.
16. Active medical devices that generate ionizing radiation and are intended for radiological diagnostics and therapy, including medical devices for monitoring or controlling such products, belong to class 2b. 
17. Active medical devices intended for the administration of medicinal products, physiological fluids or other substances into the human body and/or for their removal from the body belong to class 2a. If such administration (removal) can lead to potential hazard for a patient, considering the type of relevant substances, body parts and methods of use, the medical devices belong to class 2b.
18. Active medical devices, for which paragraphs 14 through 17 of these Rules do not apply, belong to class 1.
5. Particulars of classification of certain medical devices

19. Medical devices containing substances that, when used independently, can be considered as medicines, as well as products derived from human blood or plasma, and which affect the human body in addition to the effects of the medical device, belong to class 3.
20. Medical devices designed to control conception or to protect against sexually transmitted diseases belong to class 2b. If the specified medical devices are implantable or invasive medical devices for long-term use, they belong to class 3.
21. Medical devices intended for disinfection of invasive medical devices, as well as for cleaning, washing, disinfecting, hydrating contact lenses, belong to class 2b. Other medical devices intended for disinfection or sterilization of medical devices belong to class 2a.
22. Medical devices intended for recording images obtained from X-ray, magnetic resonance, ultrasound and other diagnostic devices belong to class 2a.
23. Medical devices manufactured using dead tissues or animal cells or their derivatives belong to class 3. If the medical devices are intended to contact with intact skin only, they belong to class 1.
24. Packages (polymer containers) for blood belong to class 2b.
25. Medical devices containing nanomaterial belong to class 3. If the nanomaterial is in an isolated or bound state, which excludes its entry into the patient's or user's body, such a medical device belongs to class 1.
26. Medical devices intended for apheresis, including kits, connectors and solutions, belong to class 3.
27. If a medical device is intended to be used in combination with another medical device, then the provisions of these Rules apply to each medical device separately.
28. If, considering the information provided by the manufacturer with regard to a medical device, several paragraphs of these Rules are applicable, the paragraph, according to which the medical device class corresponding to the highest degree of potential risk of use is established, shall be applied.
6. Algorithm of the 
medical device classification
depending on potential risk of use
29. When classifying medical devices, the applicability of all provisions of these Rules relating to the classification of medical devices shall be evaluated.
30. The medical devices classification (except for the IVD medical devices and software being medical device), depending on the potential risk of use, shall be performed in accordance with the algorithm described in Annex No. 1. 
III. Classification of IVD medical devices
, except for software being medical device
1. Classes of the IVD medical devices 
depending on the potential risk of use
31. The IVD medical devices depending on potential risk of use divide into 4 classes: 1, 2a, 2b and 3. 
Each IVD medical device may be assigned to one class only.
The IVD medical devices shall be classified basing on the following:
Class 1 includes IVD medical devices with low potential risk of use for an individual and low potential risk of use for public health;
Class 2a includes IVD medical devices with medium potential risk of use for an individual and low potential risk of use for public health;
Class 2b includes IVD medical devices with high potential risk of use for an individual and/or medium potential risk of use for public health;
Class 3 includes IVD medical devices with high potential risk of use for an individual and high potential risk of use for public health;
32. When classifying an IVD medical device, the following criteria shall be considered:
a) purpose of the IVD medical device and indications for its use specified by the manufacturer;
b) technical, scientific or medical knowledge of the intended user (non-professional or professional);
c) importance of information for diagnosis (single determinant or one of several ones), considering presence of the disease signs and symptoms, condition or other clinical information that a clinician may be guided by;
d) impact of the result (true or false) on the health of an individual and/or on public health.
2. Particulars of the IVD medical device classification depending on the potential risk of use
33. If, considering the information provided by the manufacturer with regard to an IVD medical device, several paragraphs of these Rules are applicable, the paragraph, according to which the IVD medical device class corresponding to the highest degree of potential risk of use is established, shall be applied.
34. Calibration and control materials with quantitatively and qualitatively specified values belong to the same class as the IVD reagent (reagent kit), for use with which they are intended. 
3. Classification of IVD medical devices
35. The IVD medical devices designed to detect pathogens in blood, blood components, blood derivatives, cells, tissues or organs in order to assess the possibility of their transfusion or transplantation, as well as the IVD medical devices designed to detect pathogens that cause long-term disability or death, with high or presumed high risk of spread, are classified in class 3. 
36. The IVD medical devices used to determine blood groups or to determine the group incompatibility of maternal and fetal blood or tissue types in order to guarantee the immunological compatibility of blood, blood components, cells, tissues or organs intended for transfusion or transplantation, belong to class 2b, except for those used to determine antigens or antibodies of the AB0 system (A (AB01), B (AB02), AB (AB03)), Rh system (Rh1 (D), Rh2 (C), Rh3 (E), Rh4 (c), Rh5 (e) and weak or partial Rh (D)), Kell systems (Kell (K)), Kidd systems (JK1 (Jka), JK2 (Jkb)) and Duffy systems (FY1 (Fya), FY2 (Fyb)) that belong to class 3. 
37. The IVD medical devices belong to class 2b if they are intended for:
a) detection of pathogens of sexually transmitted diseases;
b) detection of pathogens with moderate risk of spread in cerebrospinal fluid or blood; 
c) detection of pathogens if there is a risk that erroneous result of the test may cause death or disability of a patient being examined, the fetus or embryo or the patient's children; 
d) women screening during pregnancy to determine their immune status in relation to infectious diseases;
e) determining the infectious disease status or immune status in the presence of the risk that an erroneous result will lead to a therapeutic decision causing disability or inevitable occurrence of dangerous situation for life of a patient or his/her children;
f) screening, diagnosis or staging of cancer; 
g) genetic tests; 
h) monitoring levels of drugs, substances or biological components when there is a risk that an incorrect result will lead to a therapeutic decision causing dangerous situation for life of a patient or his/her children;
i) assessment of condition of the patients suffering from life-threatening disorders or diseases;
j) screening for congenital diseases of the fetus or embryo;
k) screening for congenital diseases in newborns, when there is a risk that undetected and untreated diseases could lead to life-threatening situations or disability;
l) screening to select patients for selective therapy and management as accompanying diagnosis. 
38. The IVD medical devices intended for use by non-professional users belong to class 2b. If the analysis result obtained using the indicated IVD medical devices does not have critical medical status or is preliminary and requires subsequent comparison with the relevant laboratory tests, such medical devices belong to class 2a. If the IVD medical devices are intended for the purposes specified in paragraphs 35 and 36 of these Rules, such medical devices belong to class 3.
39. The IVD medical devices without measuring function, due to their objective properties may be used as general laboratory ones, but have special characteristics, in accordance with which they are intended by the manufacturer for use in IVD procedures (without indicating specific types of laboratory tests (analytes)), as well as containers for collecting and storing bioassay samples, belong to class 1.
40. The IVD medical devices, in respect of which paragraphs 35 through 39 of these Rules do not apply, including for the diagnosis of rare pathogens in the absence of the risk that erroneous result may cause disability or death of a patient, control materials without quantitatively or qualitatively specified values, analytical instruments (analyzers) with measuring function with non-fixed list of laboratory tests performed, which depends on the reagent kits (test systems) used, belong to class 2a. 
The interdependence of the instrument and the reagents used generally precludes evaluation of the instrument separately, but this does not affect its classification as class 2a.
4. Algorithm of classification
 of IVD medical devices
depending on potential risk of use
41. The IVD medical devices classification (except for software being medical device), depending on the potential risk of use, shall be performed in accordance with the algorithm described in Annex No. 2.
42. Examples of the IVD medical devices classification (except for software being medical device), depending on the potential risk of use, are given in Appendix No. 3.
IV. Classification of software
 being medical device
1. Classes of software being medical device, depending on the potential risk of use


43. Software being medical device (hereinafter referred to as software), depending on the potential risk of use, is divided into 4 classes: 1, 2a, 2b and 3.
Each software may be assigned to one class only.
The software shall be classified on the basis of the following:
class 1 includes software with low potential risk of use;
class 2a includes software with medium potential risk of use;
class 2b includes software with increased potential risk of use;
class 3 includes software with high potential risk of use;
The potential risk of use class is assigned to software in accordance with this paragraph, regardless of the class of potential risk of use of the medical device in combination with which it is used. 
2. Software classification
44. The software classification (except for the software using artificial intelligence technologies) is performed on the basis of the type of information provided, based on assessment of the impact of the software operation results on taking medical decision and the software use conditions, and is given in the table.
Table
	Software use conditions depending on its purpose
	Type of information depending on the impact of the obtained software results on the medical decision

	
	first type of information
	second type of information
	third type of information

	category A
	3
	2b
	2а

	category B
	2b
	2а
	1

	category C
	2а
	1
	1


45. Software using artificial intelligence technologies belongs to class 3.
46. The assignment of the software use conditions depending on its purpose to one of the categories is carried out in accordance with the following criteria:
a) software use conditions are classified as category A if the software is intended:
to provide emergency medical care (in the event of sudden acute diseases, conditions, exacerbation of chronic diseases threatening the patient's life);
for use during surgical intervention and/or complex therapy or to determine the need for them;
for diagnosis and/or treatment of disease/-es or conditions of a patient with high risk to public health, and/or with high individual risk;
b) software use conditions are classified as category B if the software is intended:
to provide urgent medical care (in case of sudden acute diseases, conditions, exacerbation of chronic diseases threatening the patient's life);
to provide medical care without surgery;
for diagnosis and treatment of disease/-es or conditions of a patient posing moderate risk to public health;
c) software use conditions are classified as category C if the software is intended:
for provision of scheduled medical care (when carrying out preventive measures, in the presence of diseases and conditions not threatening the life of a patient, not requiring provision of emergency and urgent medical care, and the delay in the provision of which for a certain time will not entail aggravation in the patient's condition, the threat to his life and health);
for provision of medical care using non-invasive methods and/or for use in long-term care;
for diagnosis or treatment of a patient's disease/-es or conditions of low public health risk.
47. The software use in conditions belonging to categories A and B is possible only by a medical worker who has undergone the necessary training in its use. The software use in conditions belonging to category C is possible both by a medical worker and by a patient or other individual as directed and/or under the control of a medical worker.
48. Assignment of information, depending on the impact of the software operation results on taking medical decision to one of the types is carried out in accordance with the following criteria:
a) the first type of information includes information that does not require clarification and/or addition to make an informed clinical/medical decision and indicates the need for immediate and timely action;
b) the second type of information includes information that requires clarification and/or addition in order to make an informed clinical/medical decision;
c) the third type of information includes information that does not indicate the need for immediate action.
3. Particulars of software classification
depending on potential risk of use
49. If, when classifying the software depending on potential risk of use, a combination of various types of information obtained using this software and/or categories of conditions in which it is used is possible, the software class corresponding to the greatest degree of potential application risk shall be established.
50. Examples of software classification depending on the potential risk of use are given in Appendix No. 4.
ANNEX No. 1
to the Rules for classification of medical devices depending on potential risk of use
ALGORITHM
classification of medical devices (except for IVD medical devices and software being medical device) depending on the potential risk of use

 

	Item number
	Structural component of the Rules*
	Question
	Answer
	Conclusion

	
	
	
	
	class 
of medical device
	shift
to item

	1
	2
	3
	4
	5
	6

	1
	paragraph 6
	Is the medical device invasive?

	yes
no
	–
–
	9
2

	2
3
	paragraph 7
paragraph 7
	Is the medical device intended for storing organs, parts of organs, or for storing or administering blood, other liquids, gases, vapors or tissues into the human body?

	yes
no
	–
–
	3
4

	
	
	Is the medical device used in combination with class 2a or higher active medical devices?


	yes
no
	2а
2а
	32
32

	4
	paragraph 8
	Is the medical device intended to change the biological or physico-chemical composition and properties of blood, tissues, cells, other physiological fluids or liquids that shall enter the human body?

	yes
no
	–
–
	5
6

	5
	paragraph 8
	Is the medical device function only in the particle filtration, centrifugation, gas or heat exchange?

	yes
no
	2а
2b
	32
32

	6
	paragraph 9
	Does the medical device come into contact with damaged skin?
	yes
no
	–
1
	7
32

	7
	subparagraph a 
of paragraph 9
	Is the medical device used as mechanical barrier, for compression or for absorption of exudates? 

	yes
no
	1
–
	32
8

	8
	subparagraphs b and c of paragraph 9

	Is the medical device primarily used for wounds that can only be healed by secondary adhesion (including medical devices primarily intended to affect the wound microenvironment)?

	yes
no
	2b
2а
	32
32

	9
	paragraph 10
	Is the invasive medical device a surgical product?

	yes
no
	–
–
	15
10

	10
	paragraph 10
	Is the invasive medical device intended to be attached to the class 2a or higher class active medical device?

	yes
no
	2а
–
	32
11

	11
	subparagraph a 
of paragraph 10
	Is the invasive medical device intended for short-term use?

	yes
no
	1
–
	32
12

	12
	subparagraph b 
of paragraph 10
	Is the invasive medical device intended for temporary use?

	yes
no
	–
–
	13
14

	13
	subparagraph b 
of paragraph 10
	Is the medical device used in the mouth up to the pharynx, in the ear canal up to the eardrum, or in the nasal cavity?


	yes
no
	1
2а
	32
32

	14
	subparagraph c 
of paragraph 10
	Is the medical device used in the oral cavity up to the pharynx, in the ear canal up to the eardrum or in the nasal cavity? Can the medical device

be absorbed by the mucosa?
	yes
no
	2а
2b
	32
32

	
	
	
	
	
	

	15
	paragraph 11
	Is the surgical invasive medical device intended for short-term use? 

	yes
no
	–
–
	16
21

	16
	subparagraph a 
of paragraph 11
	Is the medical device intended for diagnosing, observing, monitoring or correcting pathologies of the heart, central circulation or central nervous system in direct contact with organs or parts of these systems? 

	yes
no
	3
–
	32
17

	17
	subparagraph b 
of paragraph 11
	Is the surgical invasive medical device a reusable surgical instrument?

	yes
no
	1
–
	32
18

	18
	subparagraph c 
of paragraph 11
	Is the surgical invasive medical device intended to transmit energy in the form of ionizing radiation?

	yes
no
	2b
–
	32
19

	19
	subparagraph d 
of paragraph 11
	Is the surgical invasive medical device intended to produce biological effect or be completely or substantially resorbed?

	yes
no
	2b
–
	32
20

	20
	subparagraph e 
of paragraph 11
	Is the surgical invasive medical device intended for the administration of medicinal products by non-professional users?

	yes
no
	2b
–
	32
32

	21
	paragraph 12
	Is the surgical invasive medical device intended for temporary use?

	yes
no
	–
–
	22
28

	22
	subparagraph a 
of paragraph 12
	Is the surgical invasive medical device intended for temporary use for diagnosis, observation, control or correction of pathologies of the heart or central circulation in direct contact with organs and parts of these systems?

	yes
no
	3
–
	32
23

	23
	subparagraph b 
of paragraph 12
	Does the surgical invasive medical device for temporary use come into contact with the central nervous system?

	yes
no
	3
–
	32
24

	24
	subparagraph c 
of paragraph 12
	Is the surgical invasive medical device intended for temporary use to transmit energy in the form of ionizing radiation?

	yes
no
	2b
–
	32
25

	25
	subparagraph d 
of paragraph 12
	Is the surgical invasive medical device intended for temporary use in order to cause biological effect or be completely or substantially resorbed?


	yes
no
	3
–
	32
26

	26
	subparagraph e 
of paragraph 12
	Does the surgical invasive medical device for temporary use undergo chemical changes in the human body (except for the medical devices implanted in the teeth or intended for the administration of medicinal products)? 

	yes
no
	2b
2а
	27
32

	27
	subparagraph e 
of paragraph 12
	Is the surgical invasive medical device for temporary use implantable in the teeth?

	yes
no
	2а
2b
	32
32

	28
	subparagraph a 
of paragraph 13
	Is the implantable medical device or surgical invasive medical device intended for long-term use for implantation in teeth?
	yes
no
	2а
–
	32
29

	29
	subparagraph b 
of paragraph 13
	Does the implantable medical device or surgical invasive medical device for long-term use contact directly with heart, central circulation or central nervous system?

	yes
no
	3
–
	32
30

	30
	subparagraph c 
of paragraph 13
	Is the implantable medical device or surgical invasive medical device intended for long-term use in order to produce biological effect or be completely or substantially resorbed?

	yes
no
	3
–
	32
31

	31
	subparagraph d 
of paragraph 13
	Does the implantable medical device or the surgical invasive medical device for long-term use undergo chemical changes in the human body (except for the medical devices implanted in the teeth or intended for the administration of medicinal products)?

	yes
no
	3
2b
	32
32

	32
	paragraph 14
	Is the medical device active?

	yes
no
	–
–
	33
46

	33
	paragraph 14
	Is the active medical device a therapeutic product?

	yes
no
	–
–
	34
37

	34
	subparagraph a 
of paragraph 14
	Is the active therapeutic medical device intended to transfer energy to the human body or energy exchange?

	yes
no
	–
–
	35
36

	35
	subparagraph a 
of paragraph 14
	Does the energy transfer to the human body or the energy exchange with it pose potential hazard due to the medical device features of the active therapeutical product, taking into account the nature, density and location of the energy impact on body parts (including active medical devices designed to create ionizing radiation, radiation therapy)?

	yes
no
	2b
–
	44
36

	36
	subparagraph b 
of paragraph 14
	Is the active medical device intended to monitor or control class 2b active therapeutic medical devices?
	yes
no

	2b
2а

	44
44


	
	
	
	
	
	

	
	
	Is the active medical device intended to monitor or control active implantable medical devices?

	yes
no
	3
2b
	completed
44

	37
	paragraph 15
	Is the active medical device diagnostic?

	yes
no
	–
–
	38
44

	38
	subparagraph a 
of paragraph 15
	Is the active diagnostic medical device designed to transmit energy that can be absorbed by a human?

	yes
no
	–
–
	39
40

	39
	subparagraph a 
of paragraph 15
	Is the active diagnostic medical device designed to illuminate the patient's body in the visible spectrum?

	yes
no
	1
2а
	44
44

	40
	subparagraph b 
of paragraph 15
	Is the active diagnostic medical device intended to represent the distribution of radiopharmaceutical medicinal products administered to a patient?

	yes
no
	2а
–
	44
41

	41
	subparagraph c 
of paragraph 15
	Is the active diagnostic medical device intended to provide direct diagnosis or monitoring of vital body functions?

	yes
no
	–
–
	42
43

	42
	subparagraph c 
of paragraph 15
	Is the active diagnostic medical device specifically designed to monitor vital physiological parameters, changes in which could lead to immediate danger to the patient (for example, changes in heart function, respiration, or the central nervous system activity)?

	yes
no
	2b
2а
	44
44

	43
	paragraph 16
	Is the active medical device generating ionizing radiation intended for radiological diagnostics and therapy (including medical devices for monitoring or controlling such devices)?

	yes
no
	2b
–
	44
44

	44
	paragraphs 
17 and 18

	Is the active medical device intended to administer medicinal products, physiological fluids or other substances into the patient's body and/or remove them from the body?

	yes
no
	–
1
	45
46

	45
	paragraph 17
	Does the administration (removal) of medicinal products, body fluids or other substances (see item 44) present a potential hazard, taking into account the type of the substances concerned, the part of the body and the method of application?
	yes
no
	2b
2а
	46
46

	
	
	
	
	
	

	46
	paragraph 19
	Does the medical device contain substance that, when used alone, can be considered as a drug, as well as a product derived from human blood or plasma, and which affects the human body in addition to the effect of the medical device?




	yes
no
	3
–
	47
47

	47
	paragraph 19
	Is the medical device being used to control conception or to protect against sexually transmitted diseases?

	yes
no
	–
–
	48
49

	48
	paragraph 20
	Is the medical device for the conception control or protection against sexually transmitted diseases an implantable or invasive medical device for long-term use?

	yes
no
	3
2b
	49
49

	49
	paragraph 21
	Is the medical device intended for disinfection or sterilization of medical devices?

	yes
no
	–
–
	50
51

	50
	paragraph 21
	Is the medical device intended to decontaminate invasive medical devices or to disinfect, clean, rinse or hydrate contact lenses?

	yes
no
	2b
2а
	51
51

	51
	paragraph 22
	Is the medical device used to record images from x-ray, magnetic resonance, ultrasound and other diagnostic devices?
	yes
no
	2а
–
	52
52

	52
	paragraph 23
	Is the medical device manufactured using dead tissue or animal cells or their derivatives?

	yes
no
	–
–
	53
54

	53
	paragraph 23
	Is the medical device made using dead tissue or animal cells or their derivatives intended to come into contact only with the intact skin surface?

	yes
no
	1
3
	54
54

	54
	paragraph 24
	Is the medical device a bag (polymer container) for collecting and storing blood?

	yes
no
	2b
–
	completed
55

	55
	paragraph 25

	Does the medical device contain nanomaterial?
	yes
no

	–
–

	56
57


	56
	paragraph 25
	Is the nanomaterial included into the medical device in an isolated or bound state, excluding its entry into the patient's or user's body?

	yes
no
	1
3
	completed
completed

	57
	paragraph 26
	Is the medical device intended for apheresis?
	yes
no
	3
1
	completed
completed

	_________________
* Rules for classification of medical devices depending on potential risk of use
Note. To classify a medical device, it is necessary to answer the question allowing to classify the medical device to the appropriate class, and then to proceed with the item indicated next to the class designation. If several paragraphs of the Rules are applicable, the class corresponding to the highest degree of potential risk of use shall be set as the class of the medical device.



ANNEX No. 2
to the Rules for classification of medical devices depending on potential risk of use
ALGORITHM
classification of IVD medical devices (except for software being medical device) depending on the potential risk of use


	Item number
	Structural component of the Rules*
	Question
	Answer
	Conclusion

	
	
	
	
	medical device class
	shift
to item

	1
	2
	3
	4
	5
	6

	1
	paragraph 40
	Is the IVD medical device intended for a fixed list of laboratory tests?

	yes
no
	–
–
	2
3

	2
	paragraph 
35 and 36
	Is the IVD medical device intended to assess the possibility of blood transfusion or transplantation, to detect pathogens that can cause life-threatening diseases with a high risk of spread?

	yes
no
	3
–
	completed
5

	3
	paragraph 39
	Can the IVD medical device be used as a general laboratory product?


	yes
no
	–
2а
	4
completed


	4
	paragraph 39
	Does the IVD medical device have a measuring function?

	yes
no
	2а
1
	completed
completed


	5
	paragraphs 
37 and 38
	Is the IVD medical device intended to detect pathogens that can cause life-threatening diseases with limited risk of spread, or for use by a non-professional user?

	yes
no
	2b
2а
	completed
completed



________________
* Rules for classification of medical devices depending on potential risk of use
ANNEX No. 3
to the Rules for classification of medical devices depending on potential risk of use
EXAMPLES
of classification of IVD medical devices (except for software being medical device) depending on the potential risk of use

	Class 
	IVD Medical Devices

	1
	rinsing and buffer solutions
specialized microscopes
containers for bioassay samples
mixing devices

	2а
	selective biochemical analyzers
hematology analyzers
specialized photometers
test systems, reagent kits for routine biochemical and hematological tests
dosing devices
pregnancy tests for self-testing
ovulation tests for self-testing
тест-полоски для анализа мочи

	2b
	test systems, reagent kits, calibrators and control materials for detection of pathological antibodies against erythrocytes
test systems, reagent kits, calibrators and control materials for detection of the following intrauterine infections: rubella, toxoplasmosis
test systems, reagent kits, calibrators and control materials for diagnosis of phenylketonuria
test systems, reagent kits, calibrators and control materials for the detection of cytomegalovirus, chlamydia
test systems, reagent kits, calibrators and control materials for the detection of troponin, cyclosporine, prothrombin time
HBV monitoring markers, HCV viral load, HIV viral load, HIV and HCV geno- and subtyping 
test systems, reagent kits, calibrators and control materials for detection the following HLA-groups of tissues: DR, A, B 
blood glucose analyzers for self-diagnosis, including appropriate calibrators and control materials

	3
	test systems, reagent kits, calibrators and control materials, including tests for self-testing, for determining the following blood groups: AB0 system, Rh system (C, c, D, E, e), Kell system
test systems, reagent kits, calibrators and control materials for detection and confirmation of HIV infection markers (HIV-1 and HIV-2, HTLV-I and HTLV-II, hepatitis B, C and D)
tests for detection bacterial contamination of blood components
especially dangerous infections with a high risk of spread


ANNEX No. 4
to the Rules for classification of medical devices depending on potential risk of use
EXAMPLES
of classification of software being medical device, depending on the potential risk of use
	Class
	Software being medical device

	1
	software for remote monitoring of the health status of elderly patients with comorbid chronic diseases (automatic detection of pathological changes in blood pressure and heart rate, signs of arrhythmias, followed by notification of those who monitor and/or care for the patient, according to the collected and received data in automatic mode)*

	
	software for developing an individual rehabilitation program*

	2а
	software to assist physicians in diagnose arrhythmias*

	
	software to assist physicians in planning surgical techniques (e.g., constructing and visualizing a customized 3D anatomical model based on computed tomography images used to locate catheters on the inside of the bronchial tree and in lung tissue or to place markers in soft lung tissue)*

	2b
	radiotherapy planning system (used as an aid in treatment)*

	
	colorectal cancer risk decision support software (e.g., assessing the risk of developing colorectal cancer based on the available data on a high-risk patient when developing his individual plan of preventive measures)*

	3
	stroke decision support software (differentiation between ischemic and hemorrhagic stroke based on diagnostic images in neuro intensive care units and vascular centers, including emergency care)

	
	software to assist the physician in diagnose tuberculous or viral meningitis in children (based on the analysis of cerebrospinal fluid spectroscopy data)


________________
* Except for the software using artificial intelligence technologies.”.
____________

